
ALBASAR INTERNATIONAL FOUNDATION
5/12 NNDC Quarters, BUK Road, Kano. 
Email: albasarnig@gmail.com

REQUEST FOR QUOTATION (RFQ)
Supply and Installation of Major Ophthalmic Equipment
	RFQ Reference No.
	BIF/SQ/EQP/2026/001

	Date of issue
	Monday September 28, 2026

	Closing date and time
	Monday 12th Oct 2026, 11:59pm (local time)

	Submission address / email
	5/12 BUK Road, NNDC Qtrs, Kano. albasarnig@gmail.com

	Contact person for clarifications
	Yusuf Jibril, Procurement Manager, yusufjibril01@gmail.com

	Delivery location(s)
	Makkah Specialist Eye Hospitals (Bauchi and Oyo)



1. Introduction and Scope of Supply
Albasar International Foundation invites qualified suppliers to submit quotations for the supply, delivery, and installation of the major ophthalmic equipment listed in Section 5, for use in a Secondary Eye Centre and Vision Centres.
The equipment comprises seven (7) items:
1.	Compact Ophthalmic A/B Scan Ultrasound System
2.	Computerized Non-Contact Tonometer
3.	Automated Visual Field Perimeter – Type A (Full-Field, Goldmann Bowl)
4.	Automated Visual Field Perimeter – Type B (Compact, 60° / 160°)
5.	Auto Kerato-Refractometer
6.	Slit Lamp
7.	Ophthalmic Operating Microscope (Floor-Standing)
Quantities for each item are shown in the Price Schedule (Section 6). A separate RFQ (RFQ 2) covers the remaining general, surgical, sterilisation and optical-workshop equipment; bidders may respond to either or both.
2. Instructions to Bidders
2.1 Eligibility and documents to accompany the quotation
a)	Bidders shall be legally registered suppliers of medical/ophthalmic equipment. Attach: certificate of incorporation (CAC), valid tax clearance certificate, and evidence of tax registration.
b)	Attach a valid Manufacturer's Authorisation Letter (or authorised-distributor certificate) for each item offered, confirming that the bidder is authorised to supply, install and service the product in Nigeria.
c)	Attach evidence of regulatory registration or approval of the offered products where applicable (for example NAFDAC medical device registration) and, for electrical equipment, conformity to recognised standards (CE, ISO 13485 manufacturer certification or equivalent).
d)	Attach at least three (3) references (purchase orders or completion letters) for similar supplies made within the last three (3) years.
e)	Attach the original manufacturer's technical brochure or datasheet for every item offered. Offers without technical literature may be rejected.
2.2 Preparation of the quotation
a)	Quotations shall be in English and shall be submitted using the Price Schedule in this document. Every line item must be priced; state the make, model and country of origin of each item offered.
b)	Prices shall be quoted in Naira, fixed for the validity period, and shall be inclusive of packing, freight, delivery to the delivery location(s), and (where applicable) installation, any applicable taxes shall be shown separately.
c)	Quotations shall remain valid for sixty (30) days from the closing date.
d)	The specifications in this RFQ are minimum requirements. Equivalent or superior products are acceptable, provided the bidder demonstrates equivalence in writing. Any deviation from the specification must be clearly stated in the quotation.
2.3 Submission and clarifications
a)	Quotations shall be submitted to the address on the cover page not later than the closing date and time. Late quotations will not be considered.
b)	Requests for clarification shall be sent in writing to the contact person on the cover page at least three (3) working days before the closing date.
3. Commercial and Contractual Terms
3.1	Delivery. Delivery within two (2) weeks of the Purchase Order, on a DAP basis to the delivery location(s) stated on the cover page. Partial deliveries must be agreed in advance.
3.2	Installation. Installation by manufacturer-trained engineers, including calibration and a functional acceptance test witnessed by the Purchaser.
3.3	Warranty. Minimum twelve (12) months on parts and labour from the date of signed installation and acceptance. Bidders shall state their local service arrangements, response time, and confirm that spare parts and technical support will be available for at least five (5) years.
3.4	Documentation. Supply user manuals, service manuals and installation/calibration records in English, and warranty certificates with each item.
3.5	Inspection. Goods will be inspected on delivery. Items that are damaged, incomplete, or that do not conform to the specification shall be replaced by the supplier at its own cost.
3.6	Payment. Payment of 50% on delivery and balance after installation and acceptance.
3.7	Late delivery. Late delivery may attract liquidated damages as provided in the Purchase Order.
4. Evaluation of Quotations
Quotations will be evaluated in three stages:
1.	Administrative compliance: completeness of the quotation and of the documents listed in clause 2.1.
2.	Technical compliance: conformity of each offered item to the minimum specification, supported by the manufacturer's literature. Non-compliant items will be rejected.
3.	Financial evaluation: the contract will be awarded to the bidder with the lowest evaluated total price among technically compliant bidders. Albasar may award by item or as a package, and will take the bidder's after-sales support into account.
5. Technical Specifications
The requirements below are minimum specifications. Each item shall be supplied complete with all standard accessories, power cables, software and consumables required for immediate use.
Item 1: Compact Ophthalmic A/B Scan Ultrasound System
	Parameter
	Minimum requirement

	Scope
	Supply and installation of a compact ophthalmic A/B scan ultrasound system.

	B-Scan
	15 MHz probe; 50° exploration angle; 60 mm depth; axial resolution 115 μm or better; 256 grey levels; storage of still images and video sequences.

	A-Scan / Biometry
	11 MHz probe; 0.03 mm electronic resolution; 60 mm measurement depth; contact and immersion techniques; automatic axial-length measurement and IOL calculation.

	Special feature
	Biometry directly in B-mode, including axial-length measurement at the macular level for highly myopic eyes and eyes with posterior staphyloma.

	Additional requirements
	Colour touch screen; DICOM/EMR compatibility; patient database; USB, Ethernet and HDMI connectivity; compact, portable design with weight below 4 kg.


Item 2: Computerized Non-Contact Tonometer
	Parameter
	Minimum requirement

	Scope
	Supply and installation of a computerized non-contact tonometer.

	Measuring range
	1–60 mmHg, with selectable 0.1 / 1 mmHg indication steps.

	Measuring modes
	1–30 mmHg and 1–60 mmHg.

	Display
	Wide 8.5-inch WVGA colour LCD touch-screen control panel.

	Operation
	Motorised up/down movement with smooth joystick control; Soft AirPuff measurement technology.

	Functions
	Adjusted tono value using manually entered corneal thickness data.

	Printer
	Built-in thermal printer with automatic paper cutting and easy-load mechanism.

	Connectivity
	USB and LAN/Ethernet.

	Power supply
	100–240 VAC, 50/60 Hz.


Item 3: Automated Visual Field Perimeter – Type A (Full-Field, Goldmann Bowl)
	Parameter
	Minimum requirement

	Scope
	Supply of a full-field automated perimeter complete with required software and standard accessories.

	Field and bowl
	180° full-field testing with 30 cm Goldmann bowl; mirror projection system.

	Test types
	Standard white-on-white static perimetry and true Goldmann kinetic perimetry.

	Stimulus
	Goldmann stimulus sizes I–V; stimulus intensity 0.2–10,000 asb (47 dB).

	Test speed
	Fast threshold testing (2–4 minutes).

	Programmes
	30-2, 24-2, 10-2, glaucoma, macula, peripheral, Esterman monocular/binocular, ptosis and disability testing.

	Patient monitoring
	Integrated blink control, pupil position control and automated eye tracking.

	Software and connectivity
	Progression analysis; DICOM, EMR and Ethernet connectivity.


Item 4: Automated Visual Field Perimeter – Type B (Compact, 60° / 160°)
	Parameter
	Minimum requirement

	Scope
	Supply of an automated visual field perimeter complete with control device, software and all standard accessories.

	Visual field range
	60° monocular / 160° binocular; viewing distance 25 cm.

	Stimulus
	Goldmann III stimulus; stimulus duration 200 ms; maximum stimulus intensity 10,000 asb; background illumination 31.5 asb.

	Testing
	Standard Automated Perimetry (SAP), white-on-white; ZATA Standard and ZATA Fast threshold strategies; single and multiple stimulus suprathreshold testing; Esterman Group 1 and 2; driving tests; customisable test locations.

	Test speed
	Single-stimulus suprathreshold approx. 90 sec/eye; multiple-stimulus approx. 60 sec/eye; ZATA approx. 4 min/eye.

	Fixation control
	Heijl-Krakau fixation monitoring; LED fixation target; integrated video eye monitor.

	Software and connectivity
	Patient database; progression analysis; HFA data import; DICOM; Ethernet; EMR/practice-management compatibility; automatic database backup.

	Documentation
	The original manufacturer's technical brochure should be submitted.


Item 5: Auto Kerato-Refractometer
	Parameter
	Minimum requirement

	Scope
	Supply of an auto kerato-refractometer.

	Refractometry
	Spherical range −25 D to +22 D; cylindrical range 0 D to ±10 D; axis 0°–180°; minimum measurable pupil diameter 2 mm.

	Keratometry
	Corneal curvature radius 5.00–10.00 mm in 0.01 mm steps; corneal refractive power 67.50 D–33.75 D; corneal astigmatic power 0 D to ±10 D; axis 0°–180°.

	Operation
	8.5-inch colour LCD touchscreen; smooth X-Y joystick movement; one-touch lock; built-in easy-load printer.

	Measurement technology
	Advanced rotary-prism-based measuring system providing reliable measurements through pupils as small as 2 mm.

	PD measurement
	20–85 mm in 0.5 mm steps.

	Connectivity
	USB, RS-232C and LAN, suitable for integration with computerised vision systems and EMR/EHR environments.


Item 6: Slit Lamp
	Parameter
	Minimum requirement

	Microscope
	Galilean.

	Magnification
	6.5x, 10x, 16x, 25x, 40x.

	Optics
	Apochromatic MaxLite.

	Slit length / width
	0–12 mm / 0–12 mm.

	Apertures
	0.3, 1, 3, 5, 9 and 12 mm.

	Rotation / tilt
	Rotation 0–180°; tilt 20°.

	Filters
	Five filters including yellow Wratten.

	Illumination
	6 V 20 W halogen or LED.

	Imaging
	Optional beam splitter (to be priced separately as an optional item).

	Build quality
	Sturdy, low-maintenance construction.


Item 7: Ophthalmic Operating Microscope (Floor-Standing)
	Parameter
	Minimum requirement

	Magnification changer
	3-step manual: 0.6x, 1.0x, 1.6x.

	Eyepiece
	WF 12.5x / 18 mm widefield.

	Objective
	f = 200 mm, apochromatic.

	Total magnification
	4.7x, 7.8x, 12.5x.

	Field of view
	48 mm, 28.8 mm, 18 mm.

	Binocular tube
	45° inclined, ergonomic.

	Focus system
	Motorised, 30 mm travel, waterproof.

	Light source
	Direct LED with stepped intensity control.

	Filters
	Heat-absorbing and cobalt blue.

	Red reflex
	Supported.

	Arm / stand
	Floor stand with 400 mm arm.

	Foot control
	Waterproof, IPX6.

	Build quality
	Robust metal construction, premium build.
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6. Price Schedule
Complete all columns. Prices in Naira, inclusive of delivery, installation, commissioning and training. Delivery period offered: ____ weeks. Warranty offered: ____ months.
	S/N
	Item description
	Qty
	Make / Model offered
	Country of origin
	Unit price
	Total price

	1
	Compact Ophthalmic A/B Scan Ultrasound System
	06
	
	
	
	

	2
	Computerized Non-Contact Tonometer
	05
	
	
	
	

	3
	Automated Visual Field Perimeter – Type A (Full-Field, Goldmann Bowl)
	02
	
	
	
	

	4
	Automated Visual Field Perimeter – Type B (Compact, 60° / 160°)
	02
	
	
	
	

	5
	Auto Kerato-Refractometer
	06
	
	
	
	

	6
	Slit Lamp
	08
	
	
	
	

	7
	Ophthalmic Operating Microscope (Floor-Standing)
	6
	
	
	
	

	Sub-total (excluding WHT)
	

	WHT
	

	TOTAL QUOTED PRICE
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7. Bidder's Declaration
We have examined this Request for Quotation, and we offer to supply the items priced in the Price Schedule in accordance with its specifications, terms and conditions. Our quotation is valid for thirty (30) days from the closing date. Except for the deviations listed below (if any), our offer complies fully with the specifications.
Deviations from the specification (if none, write "None"): ____________________________________________
	Company name
	

	Address
	

	Name of authorised signatory
	

	Designation
	

	Signature
	

	Date
	

	Company stamp
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